## Slide 1: Additional points
Apparatus
Current sources
How dim can an LED glow? A tale from Amgen
Lessons learned
Electronics Review

## Slide 2: Comment on apparatus
Although I have at my disposal a lot more tools, apparatus, and components than you do, I’ve tried to develop this curriculum using only the stuff that you have, except:
  My Arduino is plugged into a $13 USB current meter:
  https://www.amazon.com/MakerHawk-3-7-30V-Voltage-Multimeter-Voltmeter/dp/B07FMQZVW2
    I also tried a cheaper one that didn’t work …
  I use a better soldering iron: https://www.amazon.com/gp/product/B09DKQPDKB/ costs $43 and is quite nice (so far)
  https://www.amazon.com/Weller-WE1010NA-Digital-Soldering-Station/dp/B077JDGY1J/ is probably nicer, I’ve used old Weller irons purchased on ebay too

## Slide 3: Measuring RTh
The Wikipedia page gives a theoretical recipe that isn’t generally safe or practical
  RTh isn’t a real, fixed R, but an equivalence who’s value may likely change as the operating point of the circuit changes.
  We therefore want to measure RTh “close” to where we’ll be operating the device

## Slide 4: Making a constant current source
Let’s simulate the circuit at right, for a range of different RLs. Here’s (our first) Matlab code:
  V0 = 12; % volts
  Rs = 10e3; % ohms
  Rl = logspace(0,5)'; % ohms
  Vl = V0 * Rl ./ (Rs+Rl); % voltage divider eqn
  il = Vl ./ Rl * 1000; % mA
  plot(Rl, il)
  xlabel('Rl/ohms');
  ylabel('il/mA')
Doesn’t look like a constant current …
  What if we just change the scale to log-log
Sidetrack
#1
V0
RS
RL

## Slide 5: Making a constant current source
  set(gca, 'xscale', 'log');
  set(gca, 'yscale', 'log’);
  grid;
  % or replot it:
  loglog(Rl, il)
  xlabel('R_l/ohms');
  ylabel('i_l/mA')
  grid;
For RL << Rs, the current looks pretty constant …
Another way to think about this: i = V0/(Rs+RL) is ~ constant for RS >> RL (because RS is constant)
Sidetrack
#1
V0
RS
RL
RS dominates
RL dominates
“Corner”
RS = RL

## Slide 6: More simulations
See voltageAndCurrent.mlx, a Matlab LiveScript that explores 2 resistor series circuits in much more detail

## Slide 7: Sidetrack … How dim can an LED glow?
At Amgen I was tasked with comparing the sensitivities of several detectors of faint light
  These $300,000-$800,000 machines measure very faint light from surface proximity assays (SPA) that convert radioactivity to light, and we had to decide which ones to buy
  They wanted me to fly around the country to test these machines, carrying radioactive biological materials and doing chemistry with these highly dangerous materials
  What would you do?
  Would you be comfortable working in radioactive labs in foreign companies?
  Would you bring your own Geiger counter or trust what you find there?
  …
Long

## Slide 8: Scintillation Proximity detectors
Amgen tasked me with buying the most sensitive Scintillation Proximity detectors
These $500k machines measure faint Cerenkov light emitted by radioactive atoms
I was asked to fly around the country with radioactive materials to compare competing machines …

## Slide 9: Or find another solution
Light is light – so what if I used LEDs as standard light sources, instead of radioactivity-induced Cerenkov radiation?
A 1.5V battery doesn’t turn on most LEDs (no light), and a 12V garage-door opener battery in series with a 10MW resistor produces I=1.2mA which was too bright.
These detectors take many seconds of exposure normally, so flashing the LED at a low duty cycle (e.g., 10msec on every second) would reduce the brightness by 100 – if the detector was linear…
Convince my client and boss that this is a better solution

## Slide 10: Luminescence Reference Plate
Challenge
  Commercial luminescence standards are too bright to compare and validate imaging luminometers for SPA
Solutions considered
  Validate the detector with radioactive assay plates
  LED’s with extremely low current (pA-nA)
  Pulse the LEDs (uA at ~1e-4 duty cycle)
  Neutral density filters

## Slide 11: SPA Bead and LED spectra
Safire measured emission spectra of several LEDs by emission scan with deep UV excitation
LED spectra was independent of current (at least over Safire’s limited dynamic range)
Blue SPA bead Em ~ 400-480 nm
Red SPA bead Em ~ 612-618 nm

## Slide 12: Dim LEDs
LED’s are fairly linear current-to-light transducers when above their junction voltage
Operating near the junction voltage, the output is highly sensitive voltage, temperature, etc.
The application involves integrated imaging for ~10 sec, so pulsing at low duty cycle can avoid nonlinearities.

## Slide 13: Flashing lights
Electronics Workbench simulation (Demo version) shows a CMOS counter and 2 quad AND gates can generate PCM of 2^8 dynamic range.

## Slide 14: So can a microcontroller
Although a Basic Stamp (BS2) costs 100x more ($50), it can be programmed with less solder…
… and 35 lines of code:
BS2
Resistor
SIP
LED
‘ bs2 SPA Ref code
‘ 16 jan 03 BR
high 0
high 1
high 2
high 3
high 4
high 5
high 6
high 7
high 8
high 9
high 10
high 11
high 12
high 13
high 14
high 15
begin:
pulsout 0,30
pulsout 1,90
pulsout 2,270
pulsout 3,810
pulsout 4,2430
pulsout 5,7290
pulsout 6,21870
pulsout 7,65610
pulsout 8,21870
pulsout 9,7290
pulsout 10,2430
pulsout 11,810
pulsout 12,270
pulsout 13,90
pulsout 14,30
pulsout 15,10
pause 2000
goto begin

## Slide 15: A Basic Stamp Code Generator
These 5 Matlab lines generate the 35 lines of BS2 source code
It’s much easier to tweak 3 parameters in the Matlab code generator than 16 in the Stamp code.
high 0
high 1
high 2
high 3
high 4
high 5
high 6
high 7
high 8
high 9
high 10
high 11
high 12
high 13
high 14
high 15
begin:
pulsout 0,30
pulsout 1,90
pulsout 2,270
pulsout 3,810
pulsout 4,2430
pulsout 5,7290
pulsout 6,21870
pulsout 7,65610
pulsout 8,21870
pulsout 9,7290
pulsout 10,2430
pulsout 11,810
pulsout 12,270
pulsout 13,90
pulsout 14,30
pulsout 15,10
pause 2000
goto begin
% generator for BS2 code for LED plate
% 16 jan 03 BR
for i = 0:15,fprintf('high %d\n',i); end
fprintf('begin:\n');
for i = 0:7,fprintf('pulsout %d,%d\n',i, round(29.9*3^i)); end % blue LED
for i = 8:15,fprintf('pulsout %d,%d\n',i, round(10*3^(15-i))); end % red LED
fprintf('pause 2000\ngoto begin\n');
Picking optimal parameters for the code took a bit of trial and error
Notes: Could use python, C, or any other programming language to make a code generator … Today you’d do this with an Arduino like the $3 nano that you each have in the toolkit and we’ll use later. And with a tweak of syntax you could run thee for loops on the Nano .. But how might running the 5 lines behave differently than the 35 lines?

## Slide 16: What did we learn?
Always try to understand what is needed, apart from what is wanted or perceived as needed
Trust physics – which told me that light is light
  The detectors won’t care if it comes from LEDs or SPA if …
  … the wavelength (spectra, energy) and intensities are comparable
Verify everything empirically
  LEDs can produce very consistent, imperceptible (to a human) glow by pulsing them with nanoamps
    12V battery + 120 MW = 100nA
Sidetrack in the sidetrack … 100MW resistors are rare because if you don’t wear gloves or wash them in solvents, your skin oils can short them …
And they also have a lot of Johnson noise (Wikipedia it).

## Slide 18: Electronics Review
Key rules
  Ohm’s Law: V = IR
  Watt’s Law: P = VI = V2/R = I2R
  Kirchoff’s Voltage Law (KVL): SV=0 around any loop
  Kirchoff’s Current Law (KCL): SI=0 into or out of any node
Key concepts
  V-I curves are a useful way to characterize devices
    Thus are often called “characteristic curves”
  Diodes are electrical “check-valves” – allow current flow in one direction
    Diodes aren’t usefully characterized by “R” since the slope of the I-V curve is approximately zero or infinity almost everywhere
  Resistors turn current into voltage and turn voltage sources into current sources … (please read that sentence again …)

## Slide 19: Review
KVL  resistors in series are voltage dividers
  V1 = VoR1/(R1+R2)
Breadboards make it easier to hook parts together
  Don’t jam too-big wires into them
  They can fail
  Not good at high frequencies
Multimeters are our principle tool for quantifying anything in electronics
  They can lie to you ...
    Won’t tell you AC voltage on a DC scale and visa versa
    Fuses will blow if you measure a powerful voltage source while plugged into the mA or A sockets
  Always check continuity before measuring resistance and the range/zero before measuring a voltage …

## Slide 20: Review
We characterize electric circuits by
  their topology (loops)
  the components within them
Components are characterized by their I-V curves
  Or approximations (e.g., piecewise linear, etc.)
  Other characteristics include maximum/minimum ratings for voltage, current, power, temperature, humidity, etc.
  Exceed these ratings at your own peril – the components I-V characteristics may profoundly change ( component failure)
We abstract Maxwell’s equations with KVL and KCL
  Voltages around any loop add to zero ( )
  Currents into any node sum to zero
  Component characteristics + topology + KVL and KCL are generally a linear system of equations with a unique solution

## Slide 21: Review
Ohm’s law is the simplest possible I-V relationship, and a good approximation for resistors
Thevenin (or Norton) equivalent circuits
  Any 2-terminal component “sees” the rest of the Ohmic circuit connected to it as a voltage (or current) source in series (or parallel) with a resistance
Diodes have very non-linear I-V curves, but away from their (negative) zener breakdown voltage and forward conduction voltage “corners”, we can model them as piecewise linear (or open and short circuits)
We’ll see how other components can also be modeled Ohmically with some tricks

## Slide 22: Metacognating
We’ve all got our own power supply, multimeter, breadboard, soldering iron, tools, …
So you can stumble and bumble and explore the world of electronics at home, on your own ...
Why should you want a better DC power supply?
  What features should it have?